## Slide 1: Hemefilter 10/10/22
Madison Butler

## Slide 2: Last week
Received blood, added antibiotics for storage
Calculated PfHb levels in rat blood:
  Assay method to stay consistent to previous experiments
  Direct optical method
Brain device development meeting
Resistance data

## Slide 3: PfHb of rat blood
Method | Justification | Result (mg/dL) | CV | Discussion
Sigma assay | Colorimetric product from reaction with PfHb. Used by Anna and Ryan in previous experiments. | 177 | 23.9 * 10-3 | Reasonable figure.
Direct 540 concentration curve | Based on ASTM method. Need to order Drabkin’s solution to follow method exactly. | 358 | 2.03 * 10-3 | Seems unreasonably high. Why is it about 2x the size of sigma assay result?
Direct 540 concentration curve with correction factor = 0.5 | Curious why result was so high. | 179 | 2.03 * 10-3 | Method seems to be picking up other blood components.

## Slide 4: Sigma Hb assay
Triton/NaOH method: colorimetric product produced by reaction with reagent, absorbance measured at 400nm
Used by Anna & Ryan in past experiments
Result: 177 +/- 4.23 mg/dL
[Hb] =

## Slide 5: Blood Gas Analyzer
Measure total Hemoglobin from BGA
Lyse blood completely in water and make further water dilutions
  Measure dilutions at concentration curve at 540 nm
  Measure sample of un-lysed plasma at 540nm
  Relate plasma as a dilution of total hemoglobin, find concentration of PfHb
Sample based on ISO standard, using BGA. Not sure on precise method.
Result: 201 +/- 0.41 mg/dL
Chart: Dilution of lysed blood
| 0.01 | 5.0000000000000001E-3 | 2.5000000000000001E-3 | 1.25E-3 | 6.2500000000000001E-4 | 3.1250000000000001E-4 | 1.5625E-4 | 7.8125000000000002E-5
Absorbance | 0.19533 | 0.093333 | 0.042333 | 0.019 | 0.0048667 | -6.6667e-05 | -0.0036 | -0.0063

## Slide 6: 540nm direct method
Based on ASTM method but using different diluent that the ASTM lists as a replacement
Once panther express is set up will try with Drabkin’s solution (will be able to replicate ASTM exactly then)
Result: 358 +/- 0.73 mg/dL
Result with correction factor of 0.5: 179 +/- 0.36 mg/dL
Chart: Conc Hb mg/dL
| 250 | 200 | 150 | 100 | 50 | 25
Absorbance - blank | 0.258 | 0.2083 | 0.15583 | 0.1025 | 0.046867 | 0.022

## Slide 7: Plan for [Hb]
It’s important to be able to reliably know PfHb values of ex-vivo materials to characterize benchtop experiments accurately
It’s important to reproduce previous experiments well
Going to repeat these measurements Tuesday:
  If values return similar: sigma assay method will be used as most accurate method to stay consistent to previous experiments
  If values are a bit different: use average of assay values
  If values are very different: delay Hb under flow experiment until PfHb levels can be determined

## Slide 8: Resistance testing
Chart: R (mmHg/(mL/min))
| HF-001 | HF-002 | HF-003
Old Top Column | 9.6183 | 8.3771 | 7.9647
New Top Column | 9.375 | 9.9194 | 8.7165
Variability in data
Columns have a population average of 9.00 mmHg/(mL/min) and a coefficient of variation of 84.8 * 10-3
Haven’t investigated effect of gap size or mass beads
4/6 measurements between +/- 10% of pop, 2/6 between +/2 5% of pop
Two tailed t-test found difference between old and new column populations (p < .05)
New column produces 7.60% higher resistance values on average

## Slide 9: This week
Tuesday – repeat of PfHb measurements
Tuesday – adjustment of PfHb to 150-200 mg/dL if needed
Wednesday – Hb under flow with Anna