## Slide 1: 10/17/22 Hemefilter
Madison Butler

## Slide 2: Last week
Received ovine blood from Joe Hanke (center for preclinical studies?), heparinized
Performed [PfHb] experiments
Purchasing issues mostly resolved, purchased new reagents
Tb test for IACUC

## Slide 3: 540 direct method
Chart: 10/12/22
| 0.25 | 0.2 | 0.15 | 0.1 | 0.05 | 2.5000000000000001E-2 | 0
abs - blank | 0.2423 | 0.2047 | 0.15483 | 0.10683 | 0.0452 | 0.0247 | 0
Chart: 10/13/22
| 0.25 | 0.2 | 0.15 | 0.1 | 0.05 | 2.5000000000000001E-2 | 0
abs - blank | 0.2505 | 0.19347 | 0.14297 | 0.093933 | 0.045833 | 0.0229 | 0
Create concentration curve of hemoglobin (powderized, Sigma), in PBS (made in lab, ph 7.4)
Measure concentration curve & plasma sample (ovine) at 540 nm.
Relate plasma sample to concentration curve to find [PfHb]
Good linearity
Using PBS as diluent instead of Drabkin’s Solution – ASTM recommendation
  This is likely why the measured absorbance doesn’t match other tests

## Slide 4: Assay method
Create colorimetric product with plasma sample (ovine) and reagent, measure against DI water (background) and calibration solution to find [PfHb]
Used by Anna and Ryan previously
Hard to ensure no bubbles form
Colorimetric product using Triton/NaOH
[Hb] =

## Slide 5: 540 BGA method
Create dilution curve of lysed blood (ovine) diluted in DI water
Measure total Hb in whole blood via BGA
Measure dilution curve & plasma sample (ovine) at 540 nm. Relate plasma sample to dilution curve, find dilution, then multiply by total Hb to find [PfHb]
Based on ISO standard & Katelin’s previous work
Ok linearity
Consistent between days
Not sure how BGA finds tHb levels
Chart: 10/12/22
| 9.5238095238095229E-3 | 4.7619047619047615E-3 | 2.3809523809523807E-3 | 1.1904761904761904E-3 | 5.9523809523809518E-4 | 2.9761904761904759E-4 | 1.488095238095238E-4 | 7.4404761904761898E-5
abs - blank | 0.17787 | 0.085333 | 0.038533 | 0.0149 | 0.0022 | -0.0018333 | -0.0059667 | -0.0069333
Chart: 10/13/22
| 9.5238095238095229E-3 | 4.7619047619047615E-3 | 2.3809523809523807E-3 | 1.1904761904761904E-3 | 5.9523809523809518E-4 | 2.9761904761904759E-4 | 1.488095238095238E-4 | 7.4404761904761898E-5
abs - blank | 0.29027 | 0.1761 | 0.1001 | 0.052167 | 0.020333 | 0.0088667 | 0.0037667 | 0.0020667

## Slide 6: Comparison
Chart: PfHb Concentration (mg/dL)
| Est pfHb conc (540 method) | Est pfHb conc (assay method) | Est pfHb conc (BGA method)
10/12/2022 | 31.3985 | 19.8036 | 17.7397
10/13/2022 | 84.8753 | 46.289 | 38.2096
Method | [PfHb] 10/12/2022 | [PfHb] 10/13/2022 | % Difference between days | P value between days | P value comparing method to population 10/12/22 | P value comparing method to population 10/13/22
540 direct | 31.4 | 84.9 | 92.0 | 1.47 E-05 | 0.0760 | 82.8 E-05
Assay | 19.8 | 46.3 | 80.1 | 20.2 E-05 | 0.312 | 0.00593
BGA | 17.7 | 38.2 | 73.2 | 4.39 E-05 | 0.0549 | 22.2 E-05
Cells lysed overnight
Assay and BGA methods closely linked
P value testing between assay and BGA methods for 10/12 and 10/13 respectively: 0.39 and 7.42 E-05
Assay and BGA method possibly in agreement, second day casting doubt on that conclusion

## Slide 7: Plan for [PfHb]
Ovine blood confirmed relationships seen in rat blood examination from previous week:
  BGA and assay method are likely in agreement, although sometimes statistically different
  Direct 540 method not in agreement, consistently higher
Throwing out 540 direct method: consistently very different
Try to measure ovine blood again this week, further confirmation of relationships
If BGA and assay methods agree:
  BGA method backed by ISO standard, best bet
  Assay method used previously, very simple

## Slide 8: This week
Continue to measure PfHb, need to lock down method
Purchasing issues should be resolved, can purchase basic lab supplies to continue experiments
Empty out columns & remanufacture
  Measuring column mass before and after bead addition
    Before – primed with DI or completely dry? Completely dry probably more reproducible
  Manufacturing process order – fabrication, dry weight test, leak test, bead addition, wet weight test, gap test, resistance test